i-CAT Vision Quick Start Instructions

Accessing i-CAT Data: iCAT Vision only reads data exported from the iCAT software in DICOM3 single or multi file format. First, create an “iCAT Vision” folder using Windows Explorer or My Computer. If you are going to be running iCAT Vision on multiple computers on a network, create this directory on a network drive. If you plan to run it only on a single computer, create the directory on your C drive or desktop. In this directory, create an image root folder (named “iCAT Vision Root” in the example above) to save iCAT studies to. Also, create a watch folder (named “iCAT Vision Watch” in the example above) in the “iCAT Vision” folder.

Initial Patient Database Setup: Select tools and then setup. Browse to your image root folder, select it and click on OK. The program will close the first time you setup the image root folder and when you run it again, a list of patients (those present in the folder) will appear in the top left window of the screen. Left click on a patient’s name to open a study. To Auto Transfer patients from iCAT DICOM exports to iCAT Vision image root, select “tools” and then “setup.” Under “DICOM Import Controls,” check the box beside Enable Auto-transfer. Then select browse and navigate to the watch folder you created. Check the two options under the folder text bar and then click on the sleep and watch button. The program will close any open case and minimize. Create an action in the iCAT software to export cases to the watch folder in DICOM3 format. Make sure NO FILTERS are applied while you export cases to iCAT Vision. When a case is exported to the watch folder, it will be automatically transferred to the iCAT Vision root folder.

Note: If you want cases to be available to a series of computers on a network, create the image root folder on a network drive. Then set up the image root database on all desired network computers but put iCAT Vision into Sleep and Watch mode on only ONE computer on the network at a time.

Warning: DO NOT do anything in a patient study until the status bar located at the bottom of the window indicates that the patient data is fully loaded.

Wait until					        looks like

The Preview Screen:

Double click on this panoramic view to enter the implant screen

Double click on this sagittal view to enter the ortho screen

Patient List

Double click on this frontal view to enter the MPR screen

Click on the red or green line to switch between the maxilla and mandible

Patient Info

Double click on this axial view to enter the TMJ Planning

Click and drag these blue dots to adjust the pan map

Cursor tools:

All views have this window level cursor tool. When your cursor looks like this, click and drag up, down, left, and right to adjust brightness and contrast

Move the cursor to the left of cross section images. The pan cursor tool will appear. Click and drag to move cross sections up, down, left, and right

Move the cursor to the right of cross section images. The zoom cursor tool will appear. Click and drag up and down to zoom in and out.

Slice Controls: The following slice control bar is found in various views and positions throughout the program:

Left click here and drag mouse right and left to increase and decrease slice thickness and spacing

Left click here to select MIP or radiograph viewing

Left click here and drag to move across image

Measurements:

- To make a hounsfield unit (density) measurement, Right click on a view and select “HU Statistics.” Click, drag, and click to define an area. Statistics will appear in upper right corner. A maximum of 4 HU stats can be taken at a time in a normal view and 2 in a cross section view.

- To make a linear measurement, right click on a view and select “Distance” Click, drag, and release to draw a line. Distance will appear in upper left corner. A maximum of 9 distance measurements can be taken at a time in a normal view and 2 in a cross section view.

- Right click and select “HU Stats” or “Distance” again to turn the tool off

- Right click on measurement statistics to remove, inactivate, or activate them.

HU stats

Distance stats

Click and drag these blue dots to adjust the pan map

Double click on a cross section to enlarge it

Implant Screen: 						   The Ortho Screen:

Turn orange hash marks on and off by right

The ortho screen displays the radiographic lateral cephalomentric, left MIP, right MIP, frontal, and axial MIP views

Clicking on pan or axial view and selecting

“Turn hash marks on/off” from menu

Labels: The following labels on the images help

clarify the orientation of the anatomy:

R: Right Side

P: Posterior

b: buckle

The TMJ Planning Screen:	     The MRP Screen:

Click and drag blue circles to adjust condyle mapping. Click red circle to map perpendicular condyle

Browse up and down sagittal to locate condyle

The MRP Screen allows you to browse through the axial, sagittal, and frontal views

Saving and Loading Workups

When exiting iCAT Vision or switching patients, the program will prompt you to save your workup (includes axial map, …) To save workup, select “yes.” Select “Create new workup” to enter a new title for your workup or choose an existing workup name from list to overwrite it.

To load a workup, click on a patient name, and a “Select a workup” window will appear. Select an existing workup to load it or create a new workup.